## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASB 098-2023 1st Ed
ANSI ASB 098-2023 First Edition
Standard for Mass Spectral Analysis in Forensic Toxicology
Discipline: | Toxicology
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
MS Analysis Criteria Requirements | 4 | Section Title
General Rules | 4.1 | Requirement | MS analysis may be performed by several approaches including single-stage or multiple-stage instruments at low- or high-resolution. Data may be acquired using various MS acquisition modes (examples: full-scan, SIM, MRM, SRM) and may be used to establish evidence of analyte identification. Final analyte identification shall meet the requirements of ANSI/ASB Std 113, Standard for Identification Criteria in Forensic Toxicology, First Edition.
Criteria for Monitored Ions | 4.1.1 | Requirement | Ions monitored shall be diagnostic. Ions generated through the loss of water, in-source fragmentation, or isotopic ions are permissible when suitable alternatives are not available and method validation demonstrates selectivity. Unless chromatographic separation is achieved, intact isotopomers do not allow for structural elucidation by MS and shall not be used as diagnostic ions. Adducts may be considered as diagnostic ions when validation proves the uniqueness and utility of the adduct. For derivatized analytes, the ion(s) representing the derivatizing agent itself shall not be considered as diagnostic.
MS Parameters | 4.1.2 | Requirement | When establishing the upper limit of the scan range in full spectrum acquisitions, both isotopes and adducts should be considered. Selection of precursor ion(s) for multiple-stage MS techniques shall be of appropriate resolution for the mass analyzer used. Method parameters (e.g., scan range, monitored ions, dwell time, gas settings, tune criteria) for MS analysis shall be the same as used during method validation studies and follow the requirements of ANSI/ASB Std 036, Practices for Method Validation in Forensic Toxicology, First Edition.
Quality Assurance and Quality Control | 4.1.3 | Requirement | Method performance shall be continuously monitored with a rigorous quality assurance and quality control (QA/QC) program according to ANSI/ASB Std 054, Standard for Quality Control Program in Forensic Toxicology, First Edition. Further, all applicable instrument calibration and maintenance shall follow recommendations of the manufacturer and be properly documented. The laboratory procedures shall describe the instrument calibration and maintenance requirements.
Single-Stage Mass Analysis Using a Low-Resolution Mass Analyzer | 4.2 | Section Title
Full-Scan Acquisition using a Single-Stage Low-Resolution Mass Analyzer | 4.2.1 | Requirement | In addition to the criteria in 4.1, the following shall be met when using a single-stage low-resolution mass analyzer in full-scan mode. a) A minimum of a single diagnostic ion shall be monitored. b) When monitoring more than one diagnostic ion: 1. ratios of diagnostic ions shall agree with those calculated from a concurrently analyzed reference material given the tolerances shown in Table 1; OR 2. the spectrum shall be compared using an appropriate library search and be above a predefined match factor as demonstrated through method validation. c) No ions shall be present at a relative abundance equal to or greater than 50% of the base peak that are not also present in the reference spectrum. Deconvoluted mass spectra or background subtractions may be used to achieve this requirement but shall be documented when used.
SIM Acquisition using Single-Stage Low-Resolution Mass Analyzer | 4.2.2 | Requirement | In addition to the criteria in 4.1, a minimum of a single diagnostic ion may be monitored using a single-stage low-resolution mass analyzer in SIM mode. When monitoring more than one diagnostic ion, ratios shall agree with those calculated from a concurrently analyzed reference material given the tolerances shown in Table 1.
Multiple-Stage Mass Analysis (MS/MS or MSn) Using Low-Resolution Mass Analyzer | 4.3 | Section Title
Product Ion Scan Acquisition Using a Low-Resolution Mass Analyzer | 4.3.1 | Requirement | In addition to the criteria in 4.1, the following shall be met when using a multiple-stage low resolution mass analyzer in product ion scan mode: a) A minimum of a single diagnostic product ion shall be monitored. b) When monitoring more than one diagnostic product ion: 1) ratios of diagnostic ions shall agree with those calculated from a concurrently analyzed reference material given the tolerances shown in Table 1; OR 2) the spectrum shall be compared using an appropriate library search and be above a predefined match factor as demonstrated through method validation. c) No ions shall be present at a relative abundance equal to or greater than 50% of the base peak that are not also present in the reference spectrum. Deconvoluted mass spectra or background subtractions may be used to achieve this requirement but shall be documented when used.
MRM Acquisition Using Low-Resolution Mass Analyzer | 4.3.2 | Requirement | In addition to the criteria in 4.1, the following shall be met when using a multiple-stage low-resolution mass analyzer in MRM mode: a) A minimum of a single diagnostic product ion shall be monitored. b) When monitoring more than one diagnostic product ion (from either the same or different precursor ions), ion ratios shall agree with those calculated from a concurrently analyzed reference material given the tolerances shown in Table 1.
Single-Stage Mass Analysis Using High-Resolution Mass Analyzer | 4.4 | Requirement | In addition to the criteria in 4.1, the following shall be met when using a single-stage high-resolution mass analyzer or multi-stage high-resolution mass analyzer for single-stage mass analysis: a) The spectrum shall contain the molecular ion. b) All monitored diagnostic ions shall be within a pre-defined mass error and isotopic distribution tolerance as demonstrated through method validation. c) When monitoring more than one diagnostic ion: 1) ratios of diagnostic ions shall agree with those calculated from a concurrently analyzed reference material given the tolerances shown in Table 1; OR 2) the spectrum shall be compared using an appropriate library search and be above a predefined match factor as demonstrated through method validation.
Multiple-Stage Mass Analysis (MS/MS or MSn) Using A High-Resolution Mass Analyzer | 4.5 | Section Title
Product Ion Scan Acquisition Using a High-Resolution Mass Analyzer | 4.5.1 | Requirement | In addition to the criteria in 4.1, the following shall be met when using a multiple-stage high-resolution mass analyzer in product ion scan mode: a) A minimum of a single diagnostic product ion shall be monitored b) All monitored diagnostic ions shall be within a pre-defined mass error and isotopic distribution tolerance as demonstrated through method validation. c) When monitoring more than one diagnostic product ion: 1) ratios of diagnostic ions shall agree with those calculated from a concurrently analyzed reference material given the tolerances shown in Table 1; OR 2) the spectrum shall be compared using an appropriate library search and be above a predefined match factor as demonstrated through method validation. d) No ions should be present at a relative abundance equal to or greater than 50% of the base peak that are not also present in the reference spectrum. Deconvoluted mass spectra or background subtractions may be used to achieve this requirement but shall be documented when used.
MRM Acquisition Using High-Resolution Mass Analyzer | 4.5.2 | Requirement | In addition to the criteria in 4.1, the following shall be met when using a multiple-stage high-resolution mass analyzer in MRM mode: a) A minimum of a single diagnostic product ion shall be monitored. b) When monitoring more than one diagnostic product ion (from either the same or different precursor ions), ion ratios shall agree with those calculated from a concurrently analyzed reference material given the tolerances shown in Table 1.
Maximum Permitted Tolerances for Ion Ratios | Table 1 | Section Title
Relative Intensity (% of base peak or target ion) Greater than 50% | Table 1 | Requirement | Tolerances for Electron Impact Ionization (relative) = ± 20% Tolerances for all Other Ionization Techniques (relative) = ± 20%
Relative Intensity (% of base peak or target ion) 20 to 50% | Table 1 | Requirement | Tolerances for Electron Impact Ionization (relative) = ± 20% Tolerances for all Other Ionization Techniques (relative) = ± 25%
Relative Intensity (% of base peak or target ion) 10 to 20% | Table 1 | Requirement | Tolerances for Electron Impact Ionization (relative) = ± 20% Tolerances for all Other Ionization Techniques (relative) = ± 30%
Relative Intensity (% of base peak or target ion) Less than 10% | Table 1 | Requirement | Tolerances for Electron Impact Ionization (relative) = ± 50% Tolerances for all Other Ionization Techniques (relative) = ± 50%
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics